Technical Specifications of  Blood Transportation Boxes as approved by the Technical Committee for Specifications of ‘Equipments & Diagnostics for Blood Banks’ in its meeting held on 13th May 2008.

- Fixed two temperatures setting adjustable to +4°C and +22°C.

- Capacity to accommodate minimum 48 blood bags of 450 ml.

- External dimension of each transportation box should be within 1000 -1500mm (L)X 800-1000 mm (W)X 600- 700 mm (H)

- Chest / Vertical Cabinet with built in handles and hinged door.

- Body material should be Polyethylene (corrosion free)

- Smooth castors with stabilizers for flexibility and movement.

- Ambient working temperature -2°C to +43°C

- Main power switch.

- Microprocessor based temperature controller with integrated audio visual temperature alarm function on digital monitoring display.

- Separate selector switch for + 4°C and +22°C temperature   settings.

- Nominal DC power 12 / 24 volt and Nominal AC power 100 – 240 volt, 50/60Hz.

- Automatic AC / DC power supply selection.

- Automatic Cooling / heating operation.

- Air cooled hermetically sealed DC compressor.

- Vehicle fixation kit.

- Controlled fan cooling system for constant air/temperature   distribution.

- For storage of Blood bags, minimum two (2) wire basket should be provided in each box.

- Power alarm audio and visual by LED. Audio alarm can be switched off. Temperature audio and visual high and low alarm by LED.

- Minimum Polyurethane insulation CFC free for the cabinet and the door.